| Parameter | Unit | Microbioreactor | 24-well Plate |
| --- | --- | --- | --- |
| Chamber Dimensions |  |  |  |
| Diameter | mm | 1.63 | 15.49 |
| Culture Area | mm2 | 2.09 | 188.4 |
| Culture Area Specified | mm2 | - | 200 |
| Height | mm | 0.25 | 2.5 |
| Nominal Volume | μL | 0.52 | 500 |
| Surface-Area-to-Volume (SAV) Ratio | mm2/mm3 | 4 | 0.4 |
| Plate to Bioreactor Concentration Factor | - | 10 | 1 |
| Column Dimensions |  |  |  |
| Culture Area | mm2 | 21.4 |  |
| Height | mm | 0.25 |  |
| Volume | μL | 5.4 |  |
| Array Dimensions |  |  |  |
| Culture Area | mm2 | 579 |  |
| Height | mm | 0.25 |  |
| Volume | μL | 145 |  |
| Nominal Flow Conditions |  |  |  |
| Culture Area per Column | cm2 | 0.214 |  |
| Flowrate per Unit Area | μL/h/cm2 | 10 |  |
| Channel Flowrate | μL/h | 2.143 |  |
|  | m3/s | 5.95E-13 |  |
| Total Array Flowrate | μL/h | 57.851 |  |
| Individual Syringe Flowrate | μL/h | 9.642 |  |
| a Reynolds Number, Re |  |  |  |
| Fluid Density,  | kg/m3 | 1000 |  |
| Average Velocity, Q/A (chamber at full width) | m/s | 1.461E-06 |  |
| Average Velocity, Q/A (interconnect) | m/s | 9.523E-06 |  |
| Hydraulic Diameter, Dh (full width) | m | 4.34E-04 |  |
| Hydraulic Diameter, Dh (interconnect) | m | 2.50E-04 |  |
| Fluid Viscosity,  | Pa.s | 1.00E-03 |  |
| Average Re (chamber at full width) | - | 6.33E-04 |  |
| Average Re (interconnect) | - | 2.38E-03 |  |
| b Péclet Number, Pe |  |  |  |
| Diffusivity (glucose) | m2/s | 6.00E-10 |  |
| Diffusivity (40kDa growth factor) | m2/s | 8E-11 |  |
| Average Pe (chamber at full width) (glucose) | - | 1.06E+00 |  |
| Average Pe (chamber at full width) (GF) | - | 7.91E+00 |  |
| c Shear Stress, τ |  |  |  |
| Shear stress (chamber at full width) | Pa | 3.51E-05 |  |
| Shear stress (interconnect) | Pa | 2.29E-04 |  |

Notes

a Calculated as Re=ρQDh/mA, where density () = 1103  kg/m3; Q represents flowrate (m3/s); Dh represents hydraulic diameter; viscosity () = 10-3 Pa.s; and A represents cross-sectional area. Density and viscosity assumed as for water.

b Calculated as Pe=QDh/DABA, where DAB represents the diffusion coefficient.

c Shear stress estimated as τ=6μQ/h2w, where height (h) = 2.510-4 m; and width (w) = 1.6310-3 m for chamber at full width and 2.510-4 m for chamber interconnect.