Table S2.  Evaluation of the detection efficiency for adenovirus, coronavirus and S. pneumoniae in clinical samples

|  | Adenovirus | Coronavirus |
| --- | --- | --- |
|  | Ref© + | Ref© - | Ref © + | Ref © - |
| RPM v.1 + | 8 | 1 | 22 | 2 |
| RPM v.1 - | 0 | 415 | 7 | 393 |
| Sensitivity | 100 % | 76% |
| Specificity | 99.7 % | 99.5% |
| Overall agreement | 99.8 % | 97.8% |

Ref©: reference diagnostic methods, culture and/or PCR