S8 Table. Comparison of FoundationOne Liquid CDx with the reference assay for the detection of *PIK3CA* alterations

|  |  |  |
| --- | --- | --- |
|  | **Reference Assay** |  |
| **Positive** | **Negative** | **Not Evaluable** | **Missing** | **Total** |
| **FoundationOne Liquid CDx** | **Positive** | 165 | 20 | 2 | 1 | 188 | PPA: 89.19%[83.80%, 93.27%] |
| **Negative** | 5 | 222 | 1 | 2 | 230 | NPA: 97.80%[66.40%, 78.37%] |
| **Not Tested** | 29 | 65 | 2 | 0 | 96 |  |
| **Not Evaluable** | 6 | 35 | 0 | 0 | 41 |  |
| **Total** | 205 | 342 | 5 | 3 | 555 |  |
|  | PPA: 97.06%[93.27%, 99.04%] | NPA: 91.74%[87.52%, 94.88%] |  |  |  | OPA: 93.93%[91.17%, 96.04%] |

PPA = positive percent agreement; NPA = negative percent agreement; OPA = overall percent agreement