Supporting Information Text S1

Imprecision: C-peptide

Intra-assay imprecision (mean %CV) ranged from 3.1% (E170 method) to 8.9% (Immulite method).

Methodological Bias: C-peptide

There was considerable variation in C-peptide concentration between the three platforms. In the absence of reference methodology, bias was calculated against the all-method mean for the three platforms. The maximum C-peptide concentrations were found on the E170 platform, with a mean positive bias of 20.7% and the minimum observed with the Centaur assay demonstrating a negative bias -17.3% (Supporting Information Table S4).

Imprecision:  Insulin

Mean Intra assay CV was lowest in the Immulite 2000 assay (2.8%) and the highest CV was in the Centaur assay (9.9%). However the low level sample in the Immulite 2000 was below the limit of detection for the assay, quoted by the manufacturer as being 2.0 mU/mL (Supporting Information Table S4).

Methodological Bias: Insulin

There was less variation in the insulin assays than the C-peptide but there was still considerable variation in the measured insulin levels between the three platforms. In the absence of reference methodology, bias was calculated against the all-method mean for the three platforms. The maximum insulin concentrations were found on the Centaur, with a mean positive bias of 11.4%. Minimum concentrations were observed in the Immulite 2000 assay with a negative bias of -20.8% (Supporting Information Table S4).

Supporting Information Text S1- Imprecision and Methodological Bias of C-peptide and Insulin